The defenition of each parameter in HDF5 file

time: 2023-09-07                                          %Data Acquisition Time

description: 2D Volume Phantom with XXXX                   %Data Description, such as phantom shape

acquisition.gradient: [0.56 1.12 0.56] T/m                       % The gradient of the selection field in the direction xyz [T/m]

acquisition.coilCoefficient: [0.85 1.02 0.89]                     % The coil coefficient

acquisition.numAverages: 1                                  % Number of block averages per drive-field period

acquisition.numFrames: 2                                    % Number of available measurement frames

acquisition.numPeriods: 3                                    % Number of drive-field periods within a frame

acquisition.offsetField: 0                                     % Reserved, default = 0

acquisition.drivefield.driveFrequency: [2.4510e+04 2.6042e+04 0] Hz % Drive field frequencies [Hz]

acquisition.drivefield.cycle: 6.52800e-04 s                       % The time it takes to execute a cycle [s]

acquisition.drivefield.numChannels: 2                          % Number of drive field channels

acquisition.drivefield.phase: [0 0 0] deg                         % Applied drive field phase, reserved, default = 0 [deg]

acquisition.drivefield.strength: [4 11.2 0] mT                     % Applied drive field strength [mT]

acquisition.drivefield.waveform: [‘sine’, ‘sine’, ‘none’]             % Applied drive field waveform type, such as sine, triangle, none.

acquisition.receiver.bandwidth: 1.25e+06 Hz                     % Bandwidth, default defined as half of the sampling frequency Fs [Hz]

acquisition.receiver.dataConversionFactor: 1                     % Reserved, default = 1

acquisition.receiver.numChannels: 1                            % Number of receive channels

acquisition.receiver.unit: V                                    % SI unit of the measured quantity, usually Voltage

measurement.data: [93 * 1 * 3 * 2]                             %For measurement data, the dimensions are [numHarmonic, numChannels, numPeriods, numFrames]; For the system matrix, the dimensions are [numFrames, numHarmonic, numFrames]

measurement.isBackgroundFrame: [0, 1]                        % Mask indicating for each of the frames if it is a background measurement or not

measurement.isFourierTransformed: 1                          % Flag, if the data is stored in frequency space

measurement.isFrequencySelection: 1                          % Flag, if frequencies has been selected and stored

measurement.frequencySelection: [93 * 1] Hz                    % Indices of selected frequency components [Hz]

scanner.faciliy: Beihang University                            % Facility where the MPI scanner is installed

scanner.manufacturer: Beihang University                      % Scanner manufacturer

scanner.name: 2D Narrowband MPI System                     % Scanner name

scanner.operator: Zhong Hangyu                              % User who operates the MPI scanner

scanner.topology: FFP                                       % Scanner topology (e.g. FFP, FFL, MPS)

tracer.concentration: 0.05 mg/mL                              % Concentration of solute [mg/mL]

tracer.volume: 1e-04 L                                      % Total volume of applied tracer [L]

tracer.injectionTime: 2023-09-07                              % Time at which tracer injection started

tracer.name: Perimag                                        % Name of tracer used in experiment

tracer.solute: Fe                                            % Solute, e.g. Fe

tracer.vendor: XXX                                         % Name of tracer supplier

% If the data is stored in a system matrix, it also needs to include the following content:

calibration.fieldOfView: [0.038 0.038 0.019] m                  % Field of view of the system matrix [m]

calibration.fieldOfViewCenter: [0 0 0] m                       % Center of the system matrix [m]

calibration.method: robot                                    % Method used to obtain calibration data

calibration.order: xyz                                       % Ordering of the dimensions, default is xyz

calibration.size: [19 19 19]                                  % Number of voxels in each dimension

calibration.snr: 0                                          % Reserved, default = 0

calibration.step: [0.0005 0.0005 0.001] m                      % Size of each voxel [m]